**Barnstead GenPure Pro Water System Pre-installation Checklist**

Thank you for purchasing a Thermo Scientific Water Purification System. To initiate the installation process, **you must complete all of the activities listed below**, checking each box as you confirm its accuracy, then sign, date and email the completed form to ServiceSupport.LED.Asheville@thermofisher.com or fax 888-618- 2682 before installation can be performed.

Upon receipt of your completed form, a technician will contact you to arrange a convenient time for the service representative to install your new water purification system.

**Model:**

**[ ]**  **50131956 – GenPure Pro Standard [ ]**  **50131954 – GenPure Pro UF**

**[ ]**  **50131950 – GenPure Pro UV/UF [ ]**  **50131952 – GenPure Pro UV**

**[ ]**  **50131948 – GenPure Pro UV-TOC [ ]**  **50131922 – GenPure Pro UV/UF – TOC**

**Checklist:**

**[ ]  Unpacking**: Ensure that the unit has been removed from its original packaging and is positioned at the point of final installation. Be sure to keep any accessories purchased with the unit.

**[ ]  Inlet Power**: A 120 volt dedicated power supply must be accessible within 5 feet of the intended point of installation.

**[ ]**  **Inlet Water**: A potable water supply that has been pretreated by Reverse Osmosis + Ion Exchange, Deionization, or Distillation. Quality must be at least ASTM type II. The water supply should be fitted with a 3/4 inch male pipe thread connector and proper shut­off valve which must be accessible within 5 feet of the intended point of installation.

**[ ]**  **Inlet Water Temperature:** 2-35°C

**[ ]**  **Inlet Water Pressure**: The inlet water supply must have a head pressure between 1.4 – 87 PSI (0.1 - 6 Bar). Gravity feed is possible.

**[ ]**  **Installation Location**: Unit can be installed on a bench top or the wall capable of supporting approximately 60 lbs. System must be set up in its final location. If the Genpure is to be wall mounted, the system must be secured to wall with hardware delivered with the unit. This must be completed by customer’s facilities personnel prior to arrival of the installation technician.

**[ ]  Drain**: An atmospheric drain must be available within 5 feet of the final mounting location.

**[ ]**  **Accessories/Cartridges:** The Ultra Filter (if applicable) and UV bulb (if applicable) are delivered already installed in the unit. All other accessories will be inside the shipping box or behind the systems removable front cover. You must ensure that all cartridges/accessories are onsite when the technician arrives. Please check off the accessories you received:

**[ ]  09.2005 – Filter Cartridge [ ]  09.1003 – Sterile Filter Capsule 0.2µm**

**[ ]  18.0036 – PE Hose 8 x 1mm 2m [ ]  25.0075 – Feed Water Connection Kit**

**[ ]  50129630 – Power Pack [ ]  21.0057 – Wall Mounting Screws**

**[ ]  50132200 – Connecting Cord (USA) (power cord)**

**[ ]  21.1006 – Universal Adapter (for power supply)**

**[ ]  21.1007 – Universal Holder (for power supply)**

**[ ]  50131947 – Instruction Manual**

**Training:** After the installation is complete, the service technician will provide a short training session on proper operation and care of the GenPure Pro.

**Contact Person:** a representative of the customer’s site must be available at the beginning and end of the installation to meet with the service technician

**Requested date of installation:**

Company Name:

Company Address:

Company Phone:

Contact Person:

Email Address:

Signature:

Model & Serial Number:

Date:

EMAIL the completed form to ServiceSupport.LED.Asheville@thermofisher.com.

**NOTE:** It is the customer’s responsibility to complete the activities listed above. If the items mentioned are not available when the technician arrives and the installation cannot be completed**, the customer will be responsible for charges associated with a second service call to complete the installation, including time and travel.**